## Data
Look at the SG504's specs: |  |  |  |  | All measurements made with S6 head serial B0426664
Amplitude (at reference) within 3% of indicated amplitude |  |  |  |  | using a 011-0059-02 10x attenuator and BNC-F to SMA adapter
Flatness +/-4% of amplitude at reference frequency.
The reference is a 50KHz, and a 6MHz frequency oscillator. | [6 empty] | NOTE: 0.5V trim adjustment in SG504 needed adjustment after trim of R7
In other words, your SG504 is within +/-3% at 50KHz, and 6MHz, and is
+/-7% everywhere else.
10mV/division |  |  |  |  | 50mV/division |  |  |  |  | 100mV/division
Original head as found 0.5V |  |  |  |  | Original head as found 3V |  |  |  |  | Original head as found 5V
Dial Freq | Divisions | Dial Volts | % | dB | Dial Freq | Divisions | Dial Volts | % | dB | Dial Freq | Divisions | Dial Volts | % | dB
50kHz | 5 | 0.5 | 0% | 0 | 50kHz | 6 | 3.04 | 0% | 0 | 50kHz | 5 | 5 | 0% | 0
250MHz | 5 | 0.5 | 0% | 0 | 250MHz | 5.95 | 3.04 | -0.8333% | -0.072686 | 250MHz | 5 | 5 | 0% | 0
300MHz | 5 | 0.5 | 0% | 0 | 300MHz | 5.9 | 3.04 | -1.667% | -0.14598 | 300MHz | 4.95 | 5 | -1% | -0.087296
400MHz | 4.95 | 0.5 | -1% | -0.087296 | 400MHz | 5.9 | 3.04 | -1.667% | -0.14598 | 400MHz | 4.95 | 5 | -1% | -0.087296
500MHz | 5 | 0.5 | 0% | 0 | 500MHz | 5.9 | 3.04 | -1.667% | -0.14598 | 500MHz | 4.95 | 5 | -1% | -0.087296
550MHz | 5 | 0.5 | 0% | 0 | 550MHz | 5.85 | 3.04 | -2.5% | -0.21991 | 550MHz | 4.9 | 5 | -2% | -0.17548
495MHz | 4.95 | 0.5 | -1% | -0.087296 | 495MHz | 5.85 | 3.04 | -2.5% | -0.21991 | 495MHz | 5 | 5 | 0% | 0
600MHz | 4.95 | 0.5 | -1% | -0.087296 | 600MHz | 5.9 | 3.04 | -1.667% | -0.14598 | 600MHz | 4.95 | 5 | -1% | -0.087296
700MHz | 5 | 0.5 | 0% | 0 | 700MHz | 5.85 | 3.04 | -2.5% | -0.21991 | 700MHz | 4.95 | 5 | -1% | -0.087296
800MHz | 5 | 0.5 | 0% | 0 | 800MHz | 5.85 | 3.04 | -2.5% | -0.21991 | 800MHz | 4.95 | 5 | -1% | -0.087296
900MHz | 5 | 0.5 | 0% | 0 | 900MHz | 5.9 | 3.04 | -1.667% | -0.14598 | 900MHz | 5 | 5 | 0% | 0
1000MHz | 4.95 | 0.5 | -1% | -0.087296 | 1000MHz | 5.9 | 3.04 | -1.667% | -0.14598 | 1000MHz | 5 | 5 | 0% | 0
Original head R7 adjusted 0.5V |  |  |  |  | Original head R7 adjusted 3V |  |  |  |  | Original head R7 adjusted 5V
Dial Freq | Divisions | Dial Volts | % | dB | Dial Freq | Divisions | Dial Volts | % | dB | Dial Freq | Divisions | Dial Volts | % | dB
50kHz | 5 | 0.5 | 0% | 0 | 50kHz | 6 | 3 | 0% | 0 | 50kHz | 5 | 5 | 0% | 0
250MHz | 5 | 0.5 | 0% | 0 | 250MHz | 5.9 | 3 | -1.667% | -0.14598 | 250MHz | 5 | 5 | 0% | 0
300MHz | 5 | 0.5 | 0% | 0 | 300MHz | 5.95 | 3 | -0.8333% | -0.072686 | 300MHz | 5 | 5 | 0% | 0
400MHz | 5 | 0.5 | 0% | 0 | 400MHz | 5.95 | 3 | -0.8333% | -0.072686 | 400MHz | 4.95 | 5 | -1% | -0.087296
500MHz | 5 | 0.5 | 0% | 0 | 500MHz | 5.95 | 3 | -0.8333% | -0.072686 | 500MHz | 4.95 | 5 | -1% | -0.087296
550MHz | 4.98 | 0.5 | -0.4% | -0.034813 | 550MHz | 5.9 | 3 | -1.667% | -0.14598 | 550MHz | 4.95 | 5 | -1% | -0.087296
495MHz | 4.98 | 0.5 | -0.4% | -0.034813 | 495MHz | 5.95 | 3 | -0.8333% | -0.072686 | 495MHz | 5 | 5 | 0% | 0
600MHz | 5 | 0.5 | 0% | 0 | 600MHz | 5.95 | 3 | -0.8333% | -0.072686 | 600MHz | 5 | 5 | 0% | 0
700MHz | 5 | 0.5 | 0% | 0 | 700MHz | 5.9 | 3 | -1.667% | -0.14598 | 700MHz | 4.95 | 5 | -1% | -0.087296
800MHz | 5 | 0.5 | 0% | 0 | 800MHz | 5.95 | 3 | -0.8333% | -0.072686 | 800MHz | 4.95 | 5 | -1% | -0.087296
900MHz | 5 | 0.5 | 0% | 0 | 900MHz | 5.95 | 3 | -0.8333% | -0.072686 | 900MHz | 5 | 5 | 0% | 0
1000MHz | 5 | 0.5 | 0% | 0 | 1000MHz | 6 | 3 | 0% | 0 | 1000MHz | 5 | 5 | 0% | 0
Substitute head R7 adjusted 0.5V |  |  |  |  | Substitute head R7 adjusted 3V |  |  |  |  | Substitute head R7 adjusted 5V
Dial Freq | Divisions | Dial Volts | % | dB | Dial Freq | Divisions | Dial Volts | % | dB | Dial Freq | Divisions | Dial Volts | % | dB
50kHz | 5 | 0.5 | 0% | 0 | 50kHz | 6 | 3.04 | 0% | 0 | 50kHz | 5 | 5 | 0% | 0
250MHz | 4.98 | 0.5 | -0.4% | -0.034813 | 250MHz | 6 | 3.04 | 0% | 0 | 250MHz | 5 | 5 | 0% | 0
300MHz | 5 | 0.5 | 0% | 0 | 300MHz | 6 | 3.04 | 0% | 0 | 300MHz | 4.95 | 5 | -1% | -0.087296
400MHz | 5 | 0.5 | 0% | 0 | 400MHz | 5.95 | 3.04 | -0.8333% | -0.072686 | 400MHz | 4.9 | 5 | -2% | -0.17548
500MHz | 5 | 0.5 | 0% | 0 | 500MHz | 5.95 | 3.04 | -0.8333% | -0.072686 | 500MHz | 4.9 | 5 | -2% | -0.17548
550MHz | 5 | 0.5 | 0% | 0 | 550MHz | 5.95 | 3.04 | -0.8333% | -0.072686 | 550MHz | 4.9 | 5 | -2% | -0.17548
495MHz | 4.95 | 0.5 | -1% | -0.087296 | 495MHz | 6 | 3.04 | 0% | 0 | 495MHz | 4.95 | 5 | -1% | -0.087296
600MHz | 4.95 | 0.5 | -1% | -0.087296 | 600MHz | 5.95 | 3.04 | -0.8333% | -0.072686 | 600MHz | 4.95 | 5 | -1% | -0.087296
700MHz | 5 | 0.5 | 0% | 0 | 700MHz | 5.95 | 3.04 | -0.8333% | -0.072686 | 700MHz | 4.9 | 5 | -2% | -0.17548
800MHz | 5 | 0.5 | 0% | 0 | 800MHz | 5.9 | 3.04 | -1.667% | -0.14598 | 800MHz | 4.9 | 5 | -2% | -0.17548
900MHz | 5 | 0.5 | 0% | 0 | 900MHz | 5.95 | 3.04 | -0.8333% | -0.072686 | 900MHz | 4.9 | 5 | -2% | -0.17548
1000MHz | 5 | 0.5 | 0% | 0 | 1000MHz | 6 | 3.04 | 0% | 0 | 1000MHz | 4.95 | 5 | -1% | -0.087296